購 買 (英文) 規 格 書

COMMODITY DESCRIPTION

品目番號

Item No.

關稅分類番號

HSK No.

品 目 및 規 格

Description

單位

Unit

數 量

Quantity

1

전산화단층X선촬영장치

(Whole body CT Scanner System)

SET

1

A. FEATURES

1. Volume CT technology is a revolution in medical imaging, providing huge  gains in

productivity and growing  clinical applications.

2. This system supplies exceptional computer and image processing power that

significantly enhances clinical productivity, building upon the strength of the true

interface CT technologies.

3. This system has the revolutionary detector technology, scalable data acquisition

system, and a new high performance compact gantry.

4. Full 360 degree rotation speed of sub-second available improved image quality.

B. CONSIST of SYSTEM

1. Gantry                                                        1 set

2. Patient table                                                   1 set

3. Tube unit                                                          1 set

4. Generator                                                          1 set

5. Detector system                                                    1 set

6. Operator console computer system                                  1 set

7. Independent workstation system                                     1 set

8. Cardiac complete package                                          1 set

9. Direct MPR                                                        1 set

10. New iterative reconstruction                                         1 set

11. Dicom 3.0 interface                                                 1 set

12. Accessories                                                         1 set

C. SPECIFICATIONS 1. Gantry (1) Gantry aperture size (cm) (2) Minimum rotation time (sec) (3) Rear side controller	Revolution CT 80cm 0.28sec Yes
2. Patient table (1) Horizontal scannable range (metal-free, cm) (2) Table accuracy (3) Table load capacity (kg)	200cm ± 0.06mm Max. 227kg
3. Tube unit (1) Anode heat storage capacity (MHU) (2) Anode heat dissipation (KHU/min) (3) Number of focal spot (4) Focal spot size (mm)	5.5MHU 3,000KHU/min 3 Small 0.7 x 1.0mm Large 1.2 x 1.6mm Large 1.2 x 2.0mm
4. Generator (1) Maximum output power (2) Minimum kVp (2) Maximum kVp (3) Maximum mA for single source	100kW 70kVp 140kVp 740mA
5. Detector system (1) Detector name (2) Detector material (3) Max. scan slice per rotation (4) Number of detector rows (5) Z-axis coverage per rotation	Gemstone™ Clarity Detector Gemstone™ 512 256 160 mm
6. Operator Console Computer system (1) Hardware - Main computer - CPU - Memory - Hard disk - Monitor - Image archiving - Multi-tasking console: scaning data acquisition, reconstruction, filming, transferring simultaneously - Reconstruction time (2) Basic functions - Scan and recon control and display, analysis, processing, and management - Automatic filming of user-selected Images or series of images - Two-way bi-direction intercom system (3) Dose reduction method - Dose modulation software (4) Basic softwares for operation and scanning - Maximum intensity projection (MIP), multi-planar reformation (MPR) - 3D SSD - Volume rendering - Navigator - Bolus tracking software (5) Advanced software - Vessel analysis software - Auto bone removal software - CT brain perfusion software - Whole body stroke & tumor perfusion software	Clarity OC Dual Intel Six Core Xeon 2.66GHz 48GB 1000GB Dual 24" LCD DVD&CD-RW and USB Available 65 images/sec Available Available Available 3D Dose Modulation Available Available Available Available SmartPrep VesselIQ Xpress AutoBone Xpress CT Perfusion CT whole body Perfusion
7. Independent Workstation Multi-modality processing workstation Image data is transferred from workstation to MR, CT, and X-ray modalities. (1) Hardware - Main computer - Memory - Image storage capacity - Image storage devices: CD-RW, USB - Dual flat LCD - Data export for reporting function (2) Basic softwares - Image review; Image flip, zoom, pan, inverse video, magnifying and cine (3) Advanced softwares 1) Basic softwares package - Basic 3D (MIP, MPR, MPVR) - Volume rendering software - Virtual endoscopy - Volume measure and dynamic evaluation 2) Vessel Analysis software 3) Auto bone removal software 4) Lung software 5) Lung CAD software 6) Advanced colon analysis & colon wall analysis 7) Colon CAD 8) CT Brain perfusion software 9) Whole body stroke & tumor perfusion software 10) Dental study 11) CT fusion software 12) Thoracic study 13) 4D Angio analysis software	AW Volumeshare 24GB 600GB Available Dual 24'' Flat LCD Available Available Volume Viewer Volume Viewer Volume Viewer Navigator Available VesselIQ Xpress AutoBone Xpress Advanced Lung Analysis Lung VCAR Colono Pro Colon VCAR CT Perfusion CT whole body Perfusion Dental scan Fusion Thoracic VCAR Dynamic shuttle
8. Cardiac complete application package 1) Minimum cardiac scan time 2) ECG dose modulation 3) Temporal resolution (1-segment) 4) Highest achievable temporal resolution 5) Integrated ECG device 6) Cardiac acquisition gated 7) Low dose cardiac protocol 8) Calcium scoring software 9) Cardiac analysis software 10) Cardiac function software 11) Cardiac EP software	Advanced cardiac package 0.28sec ECG dose modulation 140msec Effective 24msec (with SSF) Available Prospective, Retrospective SnapShot pulse SmartScore CardIQ Xpress CardIQ Xpress Function CardEP
9. Direct MPR	Direct MPR
10. New iterative reconstruction	ASiR-V
11. Dicom 3.0 interface	Available
12. Accessories 1) QC phantoms & phantom positioner 2) Cradle cushions set 3) Head holder package 4) Operator's table & chair 5) Operating manual & service manual (CD & book) 6) Slicker (table, foot switch) 7) Coronal positioning support 8) Table head rest 9) CT Relocation 10) CT Long Table(2000mm) for AS+ 11) Dual Injector 12) Interior 13) Airconditioner 14) Apron 15) Thyroid Protector 16) Apron Hanger 17) Lead Glove 18) LAP Relocation	1 set 1 set 1 set 1 set 1 set 1 set 1 set 1 set 1 set 1 set 1 set 1 set 2 set 10 ea 10 ea 1 ea 5 ea 1 set

D. REMARKS

1. Installation

Contractor is responsible for installation of this equipment at customer's premise.

2. Warranty

Contractor or supplier shall be responsible for free of charge after-service whenever there is

any defect or any trouble during three years after the equipment is accepted by

customer.(Include Tube & Detecter)

3. Hardware and Software

Contractor or supplier shall be available the latest hardware and software at the time of

delivery.

4. Compensation of deferment

After the contract, the supplier shall deliver the equipment within 90 days to hospital.

Otherwise, supplier shall compensate delayed amount, which will be the contract amount

multiply by deferment rate (1.5/1000) and days in arrears.